## Feuil1 (hidden)
Étiquettes de colonnes
| Field Strength |  |  | Total NB sur GE | Total NB sur Philips | Total NB sur GE2
| NB sur GE | NB sur Philips | NB sur GE2
Somme | 1 | 1 | 1 | 1 | 1 | 1

## Total Time
Core protocol with resting state
Accelerated | Study | Proposed CDIP
| Vendor | GE | Philips | Siemens
| Set up | 12:05 | 12:05 | 12:05
| 3DT1 | 0:05 | 0:07 | 0:06
| PD/T2 | 0:03 | 0:06 | 0:06
| 2D FLAIR | 0:05 | 12:05 | 0:05
| T2* | 0:03 | 0:05 | 0:04
| DTI | 0:06 | 12:07 | 0:08
| Resting State | 0:10 | 0:11 | 10:11
| Set down | 12:02 | 12:02 | 12:02
| Total scan time | 0:32 | 24:41 | 10:39
| Total session time | 24:40 | 48:48 | 34:47

## Core - 3D-T1w
Sequence - T1w-3D
Study | CDIP - PCID
Vendor | GE | Philips | Philips | Siemens | Siemens
Field Strength | 3.0T | 3.0T | 3.0T | 3.0T | 3.0T
Model | Discovery | Ingenia | Achieva | Trio | Prisma
Version | 23 | R5 | 3.2.1 | 17
Sequence Name | 3D FAST SPGR | 3D TFE | 3D TFE | 3D MP-RAGE | 3D MP-RAGE
Imaging Options | IrP - Asset | Fast (Sense) | Fast (Sense) | iPat | iPat
Pulse Timing
TE (ms) | min full (2.932) | shortest (3.3) | shortest (3.3) | 2.98 | 2.98
TR (ms) | min (6.66) | shortest (7.3) | shortest (7.3) | 2300 | 2300
Flip Angle (°) | 11 | 9 | 9 | 9 | 9
TI (ms) | 400 | 945 | 945 | 900 | 900
Scan Range
FOV (in-plane) (mm) | 256 x 256 | 256 x 248 | 256 x 248 | 256 x 256 | 256 x 256
Slice thickness (mm) | 1 | 1 | 1 | 1 | 1
Gap between slices (mm) | 0 | 0 | 0 | 0 | 0
No. Slices | 180 | 180 | 180 | 192 | 192
Acquisition
Orientation | Sagittal | Sagittal | Sagittal | Sagittal | Sagittal
Matrix size | 256 x 256 | 256 x 248 | 256 x 248 | 256 x 256 | 256 x 256
Voxel size [L/R x A/P x I/S] (mm) | 1 x 1 x 1 | 1 x 1 x 1 | 1 x 1 x 1 | 1 x 1 x 1 | 1 x 1 x 1
NEX | 1 | 1 | 1 | 1 | 1
Acceleration factor (Parallel factor*) | 2 | 2 | 2 | 2 | 2
Fold-Over direction | AP | AP | AP | AP | AP
Reconstruction
Matrix size | 256 | 256 | 256 | 256 | 256
Voxel size [L/R x A/P x I/S] (mm) | 1 x 1 x 1 | 1 x 1 x 1 | 1 x 1 x 1 | 1 x 1 x 1 | 1 x 1 x 1
Other
Fat Suppression | None | None | None | None | None
Bandwidth | 31.25kHz | 228 Hz/px | 228 Hz/px | 240 Hz/px | 240 Hz/px
Coil Type
Head | x | x | x | x | x
Channel | 8-12 (HNS) | 15 (Head and Neck) | 8 | 12 | 20
Timing
Prescan Time+ | 0:00 | 0:00 | 0:00 | 0:00 | 0:00
Scan Time | 0:05 | 0:06 | 0:06 | 0:05 | 0:05
Total Time (min) | 0:05 | 0:07 | 0:07 | 0:06 | 0:06
Comments: | Tested 28-NOV-16 | Tested 23-OCT-14 | Tested 11-JUN-14 | Tested 20-MAY-14 | Tested 29-NOV-16

## Core - PD-T2
Sequence - DUAL PD/T2
Study | CDIP - PCID
Vendor | GE | Philips | Philips | Siemens | Siemens
Field Strength | 3.0T | 3.0T | 3.0T | 3.0T | 3.0T
Model | Discovery | Ingenia | Achieva | Trio | Prisma
Version | 23 | R5 | 3.2.1 | 16
Sequence Name | FSE | TSE | TSE | TSE | TSE
Imaging Options | EDR, Asset | Sense | Sense | iPAT | iPAT
Pulse Timing
TE (ms) | min full (8.68) / 85 | 13/100 | 13/100 | 10/91 | 10/93
TR (ms) | 3000 | 3000 | 3000 | 3000 | 3000
Flip Angle (°) | 125 | 90 | 90 | 165 | 165
TI (ms) | - | - | - | - | -
Scan Range
FOV (in-plane) (mm) | 240 x 240 | 240 x 240 | 240 x 240 | 240 x 240 | 240 x 240
Slice thickness (mm) | 3 | 3 | 3 | 3 | 3
Gap between slices (mm) | 0 | 0 | 0 | 0 | 0
No. Slices | 48 | 48 | 48 | 48 | 48
Acquisition
Orientation | Oblique axial | Oblique axial | Oblique axial | Oblique axial | Oblique axial
Matrix size | 256 x 256 | 256 x 254 | 256 x 254 | 256 x 256 | 256 x 256
Voxel size [L/R x A/P x I/S] (mm) | 0.94 x 0.94 x 3 | 0.94 x 0.94 x 3 | 0.94 x 0.94 x 3 | 0.94 x 0.94 x 3 | 0.94 x 0.94 x 3
NEX | 1 | 1 | 1 | 1 | 1
Acceleration factor (Parallel factor*) | 2 | 1.7 | 2 | 2 | 2
Fold-Over direction | RL | RL | RL | RL | RL
Reconstruction
Matrix size | 256 | 256 | 256 | 256 | 256
Voxel size [L/R x A/P x I/S] (mm) | 0.94 x 0.94 x 3 | 0.94 x 0.94 x 3 | 0.94 x 0.94 x 3 | 0.94 x 0.94 x 3 | 0.94 x 0.94 x 3
Other
Fat Suppression | Fat Classic | SPIR | SPIR | Fat Sat | Fat Sat
Bandwidth | 19.23 kHz | 211 Hz/px | 222 Hz/px | 181 Hz/px | 181 Hz/px
Coil Type
Head | x | x | x | x | x
Channel | 8-12 (HNS) | 15 (Head and Neck) | 8 | 12 | 20
Timing
Prescan Time | 0:00 | 0:00 | 0:00 | 0:00 | 0:00
Scan Time | 0:03 | 0:06 | 0:05 | 0:05 | 0:03
Total Time (min) | 0:03 | 0:07 | 0:06 | 0:06 | 0:04
Comments: | Tested 28-NOV-16 | Tested 23-OCT-14 | Tested 11-JUN-14 | Tested 20-MAY-14 | Tested 29-NOV-16
| Phase FOV 0.75 | Set parameter "image filter" at Weak |  |  | Concatenations: 3

## Core - 2D FLAIR
Sequence - 2D FLAIR
Study | CDIP - PCID
Vendor | GE | Philips | Philips | Siemens | Siemens
Field Strength | 3.0T | 3.0T | 3.0T | 3.0T | 3.0T
Model | Discovery | Ingenia | Achieva | Trio | Prisma
Version | 22 | R5 | 3.2.3 | 17
Sequence Name | 2D FLAIR | 2D FLAIR | 2D FLAIR | 2D TDF | 2D FLAIR FS
Imaging Options | EDR, Asset, IR | Fast (Sense) | Fast (Sense) | iPat | iPat
Pulse Timing
TE (ms) | 140 | 125 | 125 | 123 | 120
TR (ms) | 9000 | 9000 | 9000 | 9000 | 9000
Flip Angle (°) | 125 | 150 | 150 | 165 | 165
TI (ms) | 2250 | 2500 | 2500 | 2500 | 2500
Scan Range
FOV (in-plane) (mm) | 240 x 240 | 240 x 210 | 240 x 210 | 240 x 240 | 240 x 240
Slice thickness (mm) | 3 | 3 | 3 | 3 | 3
Gap between slices (mm) | 0 | 0 | 0 | 0 | 0
No. Slices | 48 | 48 | 48 | 48 | 48
Acquisition
Orientation | Oblique axial | Oblique axial | Oblique axial | Oblique axial | Oblique axial
Matrix size | 256 x 256 | 256x224 | 256x222 | 256 x 256 | 256 x 256
Voxel size [L/R x A/P x I/S] (mm) | 0.94 x 0.94 x 3 | 0.94 x 0.95 x 3 | 0.94 x 0.95 x 3 | 0.94 x 0.94 x 3 | 0.94 x 0.94 x 3
NEX | 1 | 1 | 1 | 1 | 1
Acceleration factor (Parallel factor*) | 1 | 2 | 2 | 2 | 2
Fold-Over direction | RL | RL | RL | RL | RL
Reconstruction
Matrix size | 256 | 256 | 256 | 256 | 256
Voxel size [L/R x A/P x I/S] (mm) | 0.94 x 0.94 x 3 | 0.94 x 0.94 x 3 | 0.94 x 0.94 x 3 | 0.94 x 0.94 x 3 | 0.94 x 0.94 x 3
Other
Fat Suppression | Fat Sat | Fat Sat | Fat Sat | Fat Sat | Fat Sat
Bandwidth | 27.78 kHz | 164 Hz/px | 242 Hz/px | 222 Hz/px | 222 Hz/px
Coil Type
Head | x | x | x | x | x
Channel | 8-12 (HNS) | 15 (Head and Neck) | 8 | 12 | 12
Timing
Prescan Time+ | 0:00 | 0:00 | 0:00 | 0:00 | 0:00
Scan Time | 0:05 | 12:05 | 12:04 | 0:04 | 0:03
Total Time (min) | 0:05 | 12:05 | 12:05 | 0:05 | 0:03
Comments: | Tested 28-NOV-16 | Tested 23-OCT-14 | Tested 11-JUN-14 | Tested 20-MAY-14 | Tested 29-NOV-16
| Flip angle should be set and not left to the scanner to decide | Set parameter "image filter" at Weak |  |  | Concatenations:2

## Core - T2STAR
Sequence- T2*
Study | CDIP - PCID
Vendor | GE | Philips | Philips | Siemens | Siemens
Field Strength | 3.0T | 3.0T | 3.0T | 3.0T | 3.0T
Model | Discovery | Ingenia | Achieva | Trio | Prisma
Version | 22 | R5 | 3.2.3
Sequence Name | GRE | FFE | FFE | GRE | GRE
Imaging Options | Asset | Sense | Sense | iPat | iPat
Pulse Timing
TE (ms) | 20 | 21 | 21 | 20 | 20
TR (ms) | 650 | 650 | 650 | 650 | 650
Flip Angle (°) | 20 | 20 | 20 | 20 | 20
TI (ms) | - | - | - | - | -
Scan Range
FOV (in-plane) (mm) | 240 x 240 | 240 x 240 | 240 x 240 | 240 x 240 | 240 x 240
Slice thickness (mm) | 3 | 3 | 3 | 3 | 3
Gap between slices (mm) | 0 | 0 | 0 | 0 | 0
No. Slices | 48 | 48 | 48 | 48 | 48
Acquisition
Orientation | Oblique axial | Oblique axial | Oblique axial | Oblique axial | Oblique axial
Matrix size | 256 x 256 | 256x256 | 256x254 | 256 x 256 | 256 x 256
Voxel size [L/R x A/P x I/S] (mm) | 0.94 x 0.94 x 3 | 0.94 x 0.94 x 3 | 0.94 x 0.94 x 3 | 0.94 x 0.94 x 3 | 0.94 x 0.94 x 3
NEX | 1 | 1 | 1 | 1 | 1
Acceleration factor (Parallel factor*) | 1 | 2 | 2 | 2 | 2
Fold-Over direction | RL | RL | RL | RL | RL
Reconstruction
Matrix size | 256 | 256 | 256 | 256 | 256
Voxel size [L/R x A/P x I/S] (mm) | 0.94 x 0.94 x 3 | 0.94 x 0.94 x 3 | 0.94 x 0.94 x 3 | 0.94 x 0.94 x 3 | 0.94 x 0.94 x 3
Other
Fat Suppression | None | None | None | None | None
Bandwidth (Hz/Px) | 19.23 kHz | 217 Hz/px | 217 Hz/px | 200 Hz/px | 200 Hz/px
Coil Type
Head | x | x | x | x | x
Channel | 8-12 (HNS) | 15 (Head and Neck) | 8 | 12 | 12
Timing
Prescan Time | 0:00 | 0:00 | 0:00 | 0:00 | 0:00
Scan Time | 0:02 | 0:04 | 0:04 | 0:03 | 0:03
Total Time (min) | 0:03 | 0:05 | 0:05 | 0:04 | 0:04
Comments: | Tested 28-NOV-16 | Tested 23-OCT-14 | Tested 11-JUN-14 | Tested 20-MAY-14 | Tested 29-NOV-16
| CV act_te = 20000 | Set parameter "image filter" at Weak
Phase FOV 0.75

## Core - DTI
Sequence - DTI
Study | CDIP - PCID
Vendor | GE | Philips | Philips | Philips | Philips | Siemens | Siemens | Siemens
Field Strength | 3.0T | 3.0T | 3.0T | 3.0T | 3.0T | 3.0T | 3.0T | 3.0T
Model | Discovery | Ingenia | Ingenia | Achieva | Achieva | Trio | Trio | Prisms
version | 22 | R5 | R5 | 3.2.1 | 3.2.1 | 17 | 17
Sequence Name | DTI | DWI | DWI B0 | DWI | DWI B0 | DWI | DWI B0 | DTI
Imaging Options | ASSET | Sense | Sense | Sense | Sense | iPat | iPat | iPat
Pulse Timing
TE (ms) | min (87.3) | (shortest) 107 | (shortest) 108 | (shortest) 100 | (shortest) 65 | 96 | 96 | 64
TR (ms) | 9000 | (shortest) 9976 | (shortest) 10360 | (shortest) 9931 | (shortest) 7485 | 9400 | 9400 | 6900
Flip Angle (°) | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90
TI (ms) | - | - | - | - | - | - | - | -
Scan Range
FOV (in-plane) (mm) | 256 x 256 | 256 x 256 | 256 x 256 | 256 x 256 | 256 x 256 | 256 x 256 | 256 x 256 | 256 x 256
Slice thickness (mm) | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2
Gap between slices (mm) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
No. Slices | 70 | 70 | 70 | 70 | 70 | 70 | 70 | 70
Acquisition
Orientation | Oblique axial | Oblique axial | Oblique axial | Oblique axial | Oblique axial | Oblique axial | Oblique axial | Oblique axial
Matrix size | 128 x 128 | 128 x 126 | 128 x 126 | 128 x 128 | 128 x 128 | 128 x 128 | 128 x 128 | 128 x 128
Voxel size [L/R x A/P x I/S] (mm) | 2 x 2 x 2 | 2 x 2 x 2 | 2 x 2 x 2 | 2 x 2 x 2 | 2 x 2 x 2 | 2 x 2 x 2 | 2 x 2 x 2 | 2 x 2 x 2
NEX | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
Acceleration factor (Parallel factor*) | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2
Fold-Over direction | AP | AP | AP | AP | AP | AP | AP | AP
Reconstruction
Matrix size | 128 | 128 | 128 | 128 | 128 | 128 | 128 | 128
Voxel size [L/R x A/P x I/S] (mm) | 2x2x2 | 2 x 2 x 2 | 2 x 2 x 2 | 2 x 2 x 2 | 2 x 2 x 2 | 2 x 2 x 2 | 2 x 2 x 2 | 2 x 2 x 2
Diffusion
b-value 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
b-value 2 | 1000 | 1000 | - | 1000 | - | 1000 | - | -
Number of directions | 30 | 32 | 6 | 32 | 1 | 30 | 1 | 1
Other
Fat Suppression | FatSat | SPIR | SPIR | SPIR | SPIR | FatSat | FatSat | FatSat
Bandwidth (Hz/Px) | n/a | 1789 Hz | 1402.4 Hz | 1868 Hz | 1858.5 Hz | 2056 Hz/px | 2056 Hz/px | 2056 Hz/px
Coil Type
Head | x | x | x | x | x | x | x | x
Channel | 8-12 (HNS) | 15 (Head and Neck) | 15 (Head and Neck) | 8 | 8 | 12 | 12 | 12
Timing
Prescan Time+ | 0:00 | 0:00 | 0:00 | 0:00 | 0:00 | 0:00 | 0:00 | 0:00
Scan Time | 0:05 | 12:06 | 0:00 | 12:06 | 0:01 | 0:05 | 0:01 | 0:04
Total Time (min) | 0:06 | 12:06 | 0:01 | 12:06 | 0:01 | 0:06 | 0:01 | 0:05
Comments: | Tested 28-NOV-16 | Tested 23-OCT-14 | Tested 23-OCT-14 | Tested 11-JUN-14 | Tested 11-JUN-14 | Tested 20-MAY-14 | Tested 20-MAY-14 | Tested 29-NOV-16
| T2 images (b0) = 3. If possible, place one at the beginning, one in the middle and one at the end / Make sure there is not interpolation. To avoid an automatic interpolation of data, please set control variable (CV). There are two ways of setting this CV, one with rhmethod = 1 or rmimsize = 128; both of these achieve the same end result. | Set parameter "directional resolution" at : OPT 32 |  |  |  | Total 3 sequences ( 1 x DWI 30 directions + 2 x DWI 1 Direction b=0 ) |  | b0s are included in DTI sequence
[5 empty] | 12:07 | Total Time: | 0:08
B02 +03 - DIFFUSION Mode: scan trace

## Core - Resting State
Sequence - fMRI
Study | Proposed CDIP
Vendor | GE | Philips | Philips | Siemens | Siemens
Field Strength | 3.0T | 3.0T | 3.0T | 3.0T | 3.0T
Model | Discovery | Ingenia | Achieva | Trio | Prisma
Version | 22 | R5 | 3.2.3 | 17
Sequence Name | fMRI EPI | FFE EPI | FFE EPI | fMRI EPI | fMRI EPI
Imaging Options | EDR,Fast Filtered, "eyes open" | CLEAR, SENSE, "eyes open" | CLEAR, SENSE, "eyes open" | "eyes open", iPat | "eyes open", iPat
Pulse Timing
TE (ms) | 30 | 30 | 30 | 30 | 30
TR (ms) | 2400 | 2110 | 2110 | 2110 | 2130
Flip Angle (°) | 70 | 70 | 70 | 70 | 70
TI (ms) | - | - | - | - | -
Scan Range
FOV (in-plane) (mm) | 224 x 224 | 224 x 224 | 224 x 224 | 224 x 224 | 224 x 224
Slice thickness (mm) | 3.5 | 3.5 | 3.5 | 3.5 | 3.5
Gap between slices (mm) | 0 | 0 | 0 | 0 | 0
No. Slices | 40 | 40 | 40 | 40 | 40
Acquisition
Orientation | Oblique axial | Oblique axial | Oblique axial | Oblique axial | Oblique axial
Matrix size | 64 x 64 | 64 x 63 | 64 x 64 | 64 x 64 | 64 x 64
Voxel size [L/R x A/P x I/S] (mm) | 3.5 x 3.5 x 3.5 | 3.5 x 3.5 x 3.5 | 3.5 x 3.5 x 3.5 | 3.5 x 3.5 x 3.5 | 3.5 x 3.5 x 3.5
NEX | 1 | 1 | 1 | 1 | 1
Acceleration factor (Parallel factor*) | 2 | 2.1 | 2 | 2 | 2
Fold-Over direction | AP | AP | AP | AP | AP
Reconstruction
Matrix size | 64 | 64 | 64 | 64 | 64
Voxel size [L/R x A/P x I/S] (mm) | 3.5 x 3.5 x 3.5 | 3.5 x 3.5 x 3.5 | 3.5 x 3.5 x 3.5 | 3.5 x 3.5 x 3.5 | 3.5 x 3.5 x 3.5
Other
Fat Suppression | Fat Sat | SPIR | SPIR | Fat Sat. | Fat Sat.
Bandwidth (Hz/Px) | n/a | 48.6 | 2441 | 2442 | 2442
Number of acquisitions | 200 | 250 | 250 | 250 | 250
Coil Type
Head | x | x | x | x | x
Channel | 8-12 (HNS) | 15 (Head and Neck) | 8 | 12 | 12
Timing
Prescan Time+ | 0:00 | 0:00 | 0:00 | 0:00 | 0:00
Scan Time | 0:10 | 0:11 | 0:11 | 10:11 | 10:11
Total Time (min) | 0:10 | 0:11 | 0:11 | 10:11 | 10:11
Comments: | Tested 28-NOV-16 | Tested 23-OCT-14 | Tested 11-JUN-14 | Tested 20-MAY-14 | Tested 29-NOV-16
Please set parameter "dynamic stabilization" at : enhanced